## Slide 1
A cryogen-free low temperature scanning tunneling microscope
integrating an on-axis parabolic mirror in magnetic field
Cryogen-free STM with an on-axis parabolic mirror in magnetic field
Zerui Wang, Yuhao Rao, Suyang Wang, Zeyuan Sun, Shiwei Wu
复旦大学物理学系
Department of physics, Fudan University
STM induced luminescence on Au(111) surface with a Ag tip
Challenge for STM with Optics
STM induced Luminescence (STML)
Tip Enhanced Optical Spectroscopy and Microscopy
Photo-induced Process
STM with Optics
Spatial, Spectral and Temporal resolution
High-quality photon maps with sub-molecular resolution are relatively rare!
Background and Motivation
On-axis parabolic mirror
NA ~ 0.82
Acknowledgments
The instrumentation project was supported by the Natural Science Foundation of China (grant No. 12427807, PI: Prof. Shiwei Wu).
Comparison of photon collection modes
Tunneling electrons: 1 nA 1 × 10⁹ electron/s
Extremely low quantum yield:
10⁻⁵ ~ 10⁻³ photon/electron
Photon generated: 1×10⁴ ~ 1×10⁶ photon/s
Conclusion
STM combined with optics provides simultaneous access to atomic-scale topography, electronic structure, chemical fingerprints, and temporal dynamics. However, collection and excitation efficiency are the main instrumental challenges.
We have developed a cryogen-free low-temperature STM with an on-axis parabolic mirror (NA = 0.82) compatible with high magnetic fields (Z:3.5T).
This system overcomes the critical photon collection bottleneck in STM-optics experiments and enables high-sensitivity luminescence and TERS measurements under extreme conditions.
Our home-built experimental setup
 | Fiber | Lens | Off-axis parabolic mirror | On-axis parabolic mirror
Numerical aperture | - | 0.5 | 0.85 | 0.82
Solid angle coverage | 0.33 sr | 0.82 sr | 4.50 sr | 4.21 sr
Collection efficiency | 5.3% | 13% | 72% | 67%
Focus laser spot | - | 1.5 μm (112 μm) | 0.91 μm (2 μm) | 0.94 μm
Temperature | RT | 5 K | 100 K | Cryogen-free 1.78 K
Magnetic field | - | - | - | Z: 3.5T
Reference | Ushioda @RIKEN | Zhenchao Dong @USTC | Zagonel @UNIZAMP | This work
Scanner with an UHV superconducting magnet
STM with
Pump-Probe Technique
STML of a Single Molecule
X. H. Qiu, et al, Science. 299, 542 (2003).
Photo-induced Charge Transfer of a Single Molecule
S. W. Wu, et al, Science. 312, 1362 (2006).
Cryogen-free interface
Active vibration isolation table
Detector
Preparation chamber
&
Loadlock chamber
Supporting
frame
Mirror XYZ motor
Tip Z motor
Sample XY motor
Damping frame
Defocused Laser
Focused laser
Herringbone on Au (111)
Ga atoms on GaAs (110)
STM current noise level
STM thermal drift
~0.006 /hr (Z)
Cryogen-free low temperature
Unlimited holding time
Electric field:
TERS enhancement factor
Light collection efficiency determines whether the signal can be detected!
Laser excitation field
determines the signal enhancement!
STM performance
Base temperature：
1.781 K ± 0.004 K
Space-time-resolved imaging of quantum coherent interference
M. Garg, et al. Nat. Photon. 16, 196–202 (2022).
Raman Spectroscopy of a Single Molecule
R. Zhang, et al, Nature. 498, 82 (2013).
Y. Zhang, et al. Nature 531, 623–627 (2016).
STML spectra
2 nm
2 nm